Firmware Update procedure.

Tools:

- PC or laptop with ethernetport.

- Crossover cable or a Hub/switch and 2 Ethernet cables.

- FTP software like WARFTP and unzip /extract into a folder.

- Downloaded files from the Symmetricom web site and unzip into the C:\.

Set up:

- Install FTP application

- Connect Ethernet cable/s and hub/switch between workstation. / Laptop and XLi Ethernet connector.

- Set the subnet and default gate way the same on the XLi and the PC/laptop. Set the ip address last octet to different numbers just increment an address by 1.

- Put the do files at the root directory C:\ if not already there.

- Launch the FTP application. Proceed with configuring it.

Set up FTP

1: Click on Properties

2: Then Security,

3: Then Edit User

The screen will be like this:

4: Uncheck Disable Password, and then click on OK

5: Click on Properties

6: Then Security,

7: Then Edit User

8: Click the TAB File Access

9: Click on C:\  If it is not there you can click on add button to add c:\

10 Make sure to Uncheck the DENY box and all others should be checked as below.

11. Click the Apply then the OK buttons.

The IP address of your workstation that you have put the files on C:\ is displayed on the main screen of the WARFTP application.  Your IP address will be substituted in the examples below

The 4 Xli files released for download are:

All Xli's use the following files:

192-8000V1_84.bt

192-8001V1_104_2_2.bin

192-8002V1_104_2_2.fs

The 184-8000V64.bin file, this is the FPGA data.

On Xli's shipped after March 2005 we now allow the FPGA code to be software downloadable in the file 184-8000V64.bin. This file needs to be updated.

On Xli's purchased before March 2005, you cannot upgrade the 184-8000.bin.  FPGA less than FPGA # 184-8000V50 cannot be updated.

To verify the FPGA version,

Hit the Function 18 and proceed to the FPGA screen. If the FPGS is FPGA # 184-8000V50

Or higher the FPGA file needs to be updated. If it is not than the FPGA cannot be updated.

See XLi previous version updates for telnet instructions.

Update firmware using web interface on project  rev 1.95 and above.

Log into web page and select Burn Firmware.

Click on Browse and browse to the Boot loader firmware.

Select the file and bootloader and load as shown. Click on submit changes.

After the File is burned select Reboot in the Burn file Execution drop down.

you will see this message.

After logging in browse to the file shown below to load the software. Select software and load. click on submit changes.

After the File is burned select Reboot in the Burn file Execution drop down.

you will see this message.

After logging in browse to the file shown below to load the filesystem. Select “file system” and “load”. click on “submit changes”.

After the File is burned select Reboot in the Burn file Execution drop down.

you will see this message.

After logging in browse to the file shown below to load the FPGA file. Select FPGA and load. click on submit changes.

After the File is burned select Reboot in the Burn file Execution drop down.

you will see this message.

Verify the files are updated by logging in and checking “System Configuration”.