The 3 Xli files released for 2-3.1  download patch are:

All Xli's with the 86-8013 GPS card use the following files:

192-8000V1_105.bt

192-8001V1_105_2_1.bin

192-8002V1_105.fs

An FTP server must be used to download these files. One is available at the following link:

http://www.symmetricom.com/media/files/downloads/software/ntp-firmware/warftpd1.zip

On units purchased in and before 2005 the FPGA is not upgradable.

See XLi manual for complete upgrade instructions.

Pay attention to command and file names, if incorrect burn command are given to files,  the unit will need to be returned to the factory.

On older software versions after uploading the 192-8000V1_105.bt boot loader and then resetting the Xli with the F100 K I L L, the display will show Booting…, this is normal and do not power off the XLI,  just follow the instructions to download the next 2 files, the 192-8001V1_105_2_1.bin file and the 192-8002V1_105.fs file.

Below is a copy of a telnet session with the current firmware doing a successful firmware download.

WELCOME TO SYMMETRICOM NETWORK INTERFACE!
USER NAME: operator
PASSWORD: *****
NETWORK INTERFACE 192-8001      (c) 1998 - 2008 SYMMETRICOM
ALL RIGHTS RESERVED
LOGIN SUCCESSFUL!

>f100 bh 192.168.49.2 192-8000V1_105.bt

BURN HOST IS READY!!!

>f100 bub

OK

BURNING FILE 192-8000V1_105.bt WITH SIZE 486204 TO PARTITION:0 SECTOR:0

FILE: 486204 BYTES, PARTITION: 655360 BYTES (74% used)

SEC: 0 RE: 0

SEC: 1 RE: 0

SEC: 2 RE: 0

SEC: 3 RE: 0

SEC: 4 RE: 0

SEC: 5 RE: 0

SEC: 6 RE: 0

SEC: 7 RE: 0

FLASH SUCCESSFULLY PROGRAMMED CRC32 = 0x88CA6FF8

>f100 K I L L

OK

RESETTING THE UNIT

PLEASE WAIT...

(After this reset the display may show Booting…, this is normal and do not power off the XLI reestablish the telnet session and continue with the download process)

>f100 bh 192.168.49.2 192-8001V1_105_2_1.bin

BURN HOST IS READY!!!

>f100 bu

OK

BURNING FILE 192-8001V1_105.bin WITH SIZE 1028592 TO PARTITION:1 SECTOR:16

FILE: 1028592 BYTES, PARTITION: 5112320 BYTES (20% used)

SEC: 16 RE: 0

SEC: 17 RE: 0

SEC: 18 RE: 0

SEC: 19 RE: 0

SEC: 20 RE: 0

SEC: 21 RE: 0

SEC: 22 RE: 0

SEC: 23 RE: 0

SEC: 24 RE: 0

SEC: 25 RE: 0

SEC: 26 RE: 0

SEC: 27 RE: 0

SEC: 28 RE: 0

SEC: 29 RE: 0

SEC: 30 RE: 0

SEC: 31 RE: 0

FLASH SUCCESSFULLY PROGRAMMED CRC32 = 0x1C374ED

>f100 K I L L

OK

RESETTING THE UNIT

PLEASE WAIT...

WELCOME TO SYMMETRICOM NETWORK INTERFACE!
USER NAME: operator
PASSWORD: *****
NETWORK INTERFACE 192-8001      (c) 1998 - 2008 SYMMETRICOM
ALL RIGHTS RESERVED
LOGIN SUCCESSFUL!

>f100 bh 192.168.49.2 192-8002V1_105.fs

BURN HOST IS READY!!!

>f100 bf

OK

BURNING FILE 192-8002V1_105.fs WITH SIZE 2096640 BYTES

SEC: 94

SEC: 95

SEC: 96

SEC: 97

SEC: 98

SEC: 99

SEC: 100

SEC: 101

SEC: 102

SEC: 103

SEC: 104

SEC: 105

SEC: 106

SEC: 107

SEC: 108

SEC: 109

SEC: 110

SEC: 111

SEC: 112

SEC: 113

SEC: 114

SEC: 115

SEC: 116

SEC: 117

SEC: 118

SEC: 119

SEC: 120

SEC: 121

SEC: 122

SEC: 123

SEC: 124

SEC: 125

FILE SYSTEM FLASH BURN COMPLETED

>f100 K I L L

OK

RESETTING THE UNIT

PLEASE WAIT...

You may  now check Function 18 to verify the Project Revision. It should be PROJ REV # 2-3.1